## Slide 1: Defibrillator for Use Through a Right Minithoracotomy
Team 25:
Stephen Browning
Kathleen Kafka (presenting)
Kelly Moor
Mentors:
Dr. Damiano
Dr. Robertson
Dr. Okada
Dr. Saint

## Slide 2: Overview
Scope:
Minimize dangerous patient convulsions caused by shock
Successful and safe defibrillation
Right minithoracotomy
LIFEPAK 20e compatibility
Our Design:
Prototype design
SVC coil electrode
RV epicardial patch electrode
Connecting cables
Overview Analysis Design Materials/Manufacturing Conclusions
Notes: Need and/or specific design requirements; quick overview of chosen design concept

## Slide 3: Electrode Configuration
References
De Jongh et al.
  Finite element analysis
  Clinical trials
Kinst et al.
  Finite element analysis
Winkle et al.
  SVC coil electrode
  LV apex patch electrode
Literature search
  Similar electrode configurations
  DFT: 10-20 J
Damiano lab meeting
  Electrode placement is doable
Overview Analysis Design Materials/Manufacturing Conclusions
Notes: Specific Details of Analytic Efforts needed for design (15 pts.) (calculations for power requirements, weight, sampling speeds, BW, etc.)

## Slide 4: Electrodes: Literature, Patents, Surgeon Interviews
Transvenous SVC Coil
Surface area: 1200 mm2
  Range from similar ICD electrodes: 160-860 mm2
  Larger surface area → lower DFT
Coil length: 6 cm
  Shorter → easier to maneuver
Diameter: 2.83 mm (8.5 French)
  Must fit through central line
  Size limit: 8.5 French
Placement: SVC
  Commonly used ICD site
RV Epicardial Patch
Surface area: 16.2 cm2
  Range found in literature: 12-33 cm2
  Larger surface area → lower DFT
Length: 5.5 cm
  Rudski et al. – small RV length = 5.6 cm
  Fits on most RVs
Width: 3.5 cm
  Fits through 4 cm diameter incision
Placement: between diaphragm and RV
  Not a common site
  Used in a few case studies – Yoda et al. & Molina et al.
  Recommended by Dr. Damiano
Overview Analysis Design Materials/Manufacturing Conclusions

## Slide 5: LIFEPAK Circuit Analysis
Note: this calculation assumes a DFT of 30 J, but we expect a lower DFT of 20 J for our design
Current range reported by Kinst et al.: 6.4 – 11.9 A
Voltage reported by Kinst et al.: 615 V
Overview Analysis Design Materials/Manufacturing Conclusions

## Slide 6: SVC Electrode
Overview Analysis Design Materials/Manufacturing Conclusions
Inserted through the right internal jugular vein
  Pre-existing central line port
Placed in the SVC
  Can be checked for correct placement by feeling the SVC
Trifilar coil
  More surface area
  Flexibilty
  Redundancy

## Slide 7
Overview Analysis Design Materials/Manufacturing Conclusions
(with break)
connector
(all units are in mm)

## Slide 8: RV Epicardial Electrode
Overview Analysis Design Materials/Manufacturing Conclusions
Inserted through the right minithoracotomy incision
Placed on the inferior RV wall
  Held in place by the diaphragm
  Visual check for correct positioning
Mesh
  More surface area
  Semi-rigid
  No longer under patent protection
Notes: Specific Details of Chosen Design (15 pts.) (drawings, required parts, software overviews, safety, wiring and circuit diagrams, etc.)

## Slide 9
Overview Analysis Design Materials/Manufacturing Conclusions

## Slide 10: Safety – Risk Reduction
Electrode design – silicone “padding”
  Eliminates sharp edges
Training
  Extensions of pre-existing training
Warnings
  Sharp edges and electrical danger
  High impedance indicates incorrect placement
Manuals
  Set up and use of electrode
  Ignitable oxygen
Gloves and protective clothing
  Hazardous biological waste
  Blood borne illness
Overview Analysis Design Materials/Manufacturing Conclusions

## Slide 11: Materials, Cost, and Lead Times
Part | Description | Material | Manufacturer | Cost | Lead Time
SVC Electrode | Conductor Coil | Tantalum coated copper | McMaster Carr | $95 | 1 week
Epicardial Electrode | Conductor Mesh (1/2 ft²) | Titanium mesh (50 lines/inch) | Unique Wire Weaving Company | $120 | 2-3 days
Epicardial Electrode | Insulator Backing | Silicone sheeting | Bioplexus | $150 | 1 week
Connector Cable | Wire (100 ft) | Silicone coated copper wire | Alpha Wire | $560 | 1-2 days
Connector Cable | Polyurethane coating (36 yards) | Polyurethane tape | 3M | $1,300 | 10-15 days
Connectors | 3 connectors | polyester | Molex | $3 | 1-2 days
Overview Analysis Design Materials/Manufacturing Conclusions
Total for prototype production = $2200
Notes: Specific Material and Manufacturing definitions (15 pts.) (part prices and lead times, overview of manufacturing processes, etc.)

## Slide 12: Prototype Production
Overview Analysis Design Materials/Manufacturing Conclusions
Transvenous SVC Electrode
Strip section of silicone coated 3-conductor wire
Coil wire to defined length and diameter
Electroplate with tantalum
Attach silicone tip
Attach to cable using Molex connector
RV Epicardial Electrode
Cut silicone and titanium mesh to size
Stretch silicone over mesh
Insert Molex connector between silicone and mesh
Cure in oven (25 min @ 350 degrees F)
Crimp to cable using Molex connector

## Slide 13: Intellectual Property
Several design aspects taken from patents
  Transvenous Electrode: Intermedics, Inc. Patent
  Epicardial Electrode: Ideker Patent
  Expired patent protection
Later patents
  Specific uses
  Attachment methods (ICD)
  No patent infringement
Our IP
  Utility patent
  Temporary implantation
  Minithoracotomy
Overview Analysis Design Materials/Manufacturing Conclusions
Notes: Conclusions (10 pts.) (Did you solve the problem?, Future directions?, What did you learn? What would you do different?, Is there any IP in your Project? What? Why is it IP?)

## Slide 14: Discussion
What would we have changed?
More discussion with Dr. Damiano
Engineering mentor
Problem Solved?
Direct conduction into myocardium
  Lower energy
  Less skeletal muscle activation
  Minimizes patient movement caused by shock
Patient Safety
  No extra incisions
Fits through minithoracotomy
What did we learn?
Surgeons vs. engineers
Mechanical drawing tools
Design process
Future Directions
Application to partial sternotomy and left minithoracotomy
Patch electrode applicator
FDA Approval: Animal/Human Trials
  Optimal materials – cost versus efficacy
  Optimal placement – lower DFT
  Adjustment to electrode size and shape
Overview Analysis Design Materials/Manufacturing Conclusions
Notes: What we learned

## Slide 15: Questions?

## Slide 16
Overview Analysis Design Materials/Manufacturing Conclusions