## Slide 1: Defibrillator For Use in Minimally Invasive Surgery
Kelly Moor
Kathleen Kafka
Stephen Browning

## Slide 2: Demonstrated Need
Dramatic increase in popularity of minimally invasive surgical procedures
Smaller incision size disallows for traditional use of internal paddles to defibrillate patient
External paddles only current technology compatible with minimally invasive surgery
Result in dangerous convulsion by patient when shocked
Indicates need for redesign of defibrillation electrodes and leads to be more compatible with minimally invasive surgery

## Slide 3: Design Specifications
Internal components must fit through 4cm diameter incision size
Maximum of 20 sec should pass between fibrillation recognition and stimulus delivery
Leads must be compatible with LIFEPAK 20e defibrillator unit
Leads must cause no unintended damage to myocardium or other internal tissues
Components in contact with bodily tissue must be made of hypoallergenic material
Reusable components must be sterilizable by ethylene oxide

## Slide 4: Design Alternatives
KCl injection
  Injection of KCl into the heart
  KCl would arrest the heart by preventing action potentials from firing
  Successfully stop fibrillation, but cardiac massage often needed to restart the heart

## Slide 5: Electrode Types
Internal Patch
Manually Placed
Endocardial Lead
Left Pectoral Internal
External Patch
Abdominal Internal

## Slide 6: Internal Patch Electrodes
Electrode enclosed in flat, round, insulated patch
Placement on either left or right ventricle
Held in place by superficial sutures placed in myocardium
Left in place throughout surgery

## Slide 7: Manually Placed Electrodes
Right or left ventricular
Direct contact with ventricular epicardium
Conductive electrode attached to ergonomic, electrically insulated handle
Only applied during defibrillation
Manually held in place by surgeon during defibrillation

## Slide 8: Endocardial Electrodes
Insulated lead with electrode placed at right or left ventricular apex
Solves problem of difficult access locations by feeding lead through accessible locations of heart
Leads would enter through existing ports to prevent extra incisions
  Cardioplegia cannula
  Swan-Ganz catheter

## Slide 9: Endocardial Electrodes

## Slide 10: Left Pectoral Internal Electrodes
Subcutaneous
  Inserted through additional incision
  Held in place by a single suture and encapsulation by tissue
Thoracic cavity
  Inserted through main surgical incision
  Sutured to internal chest wall

## Slide 11: External Patch Electrode
Typical external patch placed in contact with skin
Location chosen to maximize myocardial mass between electrodes
Lower energy could be used if paired with an internal electrode
davismedical.com

## Slide 12: Abdominal Internal Electrode
Paired with two internal electrodes to reproduce common ICD configuration
Inserted through main incision site
Held in place by sutures

## Slide 13: List of Design Alternatives
Left Pectoral Internal Electrode + RV Patch Electrode
Left Pectoral External Patch Electrode + RV Patch Electrode
LV Patch Electrode + Right Pectoral External Electrode
LV and RV Patch Electrodes
Manually Placed LV + RV Patch Electrodes
LV Patch Electrode + Manually Placed RV Electrode
RV Patch Electrode + Manually Placed LV Electrode
Left Pectoral Internal + RV Manual Electrode
Abdominal Electrode + Two Internal Electrodes
LV Endocardial Lead + RV Patch
LV Endocardial Lead + Manually Placed RV Electrode
LV Endocardial Lead + External Patch
LV Patch + RV Endocardial Lead
Manually Placed LV Electrode + RV Endocardial Electrode
RV Endocardial Lead + Internal Left Pectoral Electrode
RV Endocardial Lead + External Patch
External Electrode + LV Manual Electrode
External Electrode + RV Manual Electrode
KCL Injection

## Slide 14: Analysis of Design Alternatives
Minimize Defibrillation Energy
  Assigned values based data found in literature based on clinical studies and electrode placement modeling
Minimize Added Risk to Patient
  Assigned risk values based on whether leads required blind navigation for placement, potentially resulting in perforation of the myocardium or coronary arteries
Ease of Electrode Placement
  Values based on need for blind placement, additional procedures, or sutures for attachment
Surgeon Preference
  Values assigned based on interviews with surgeons and how they believed designs would be received by fellow surgeons
Minimize Additional Procedure Time
  Assigned values based on the amount of time our device would add to the procedure

## Slide 15
Reduce Fibrillation Duration
  Values assigned based on how quickly the shock could be delivered.
  Example: Manual placement vs. electrodes left in place throughout procedure
Maximize Myocardial Mass Between Electrodes
  Studies show maximizing mass between electrodes increases probability of successful defibrillation
  Values assigned based on ventricular mass between electrodes
Patient Aesthetics
  Assigned values based on need for extra incisions or probability of causing burns on the skin
Cost
  Values based on comparison to similar electrodes on the market
Post-Operative Patient Discomfort
  Points detracted for designs requiring additional incisions or burn marks on the skin
Achievable before Deadline
  Assigned values based on expected amount of redesign of existing devices

## Slide 16: Pugh Chart
Design Criteria | Weight | Manual RV - Left Thoracic Cavity | RV Patch - Left Thoracic Cavity | Manual RV - Left Chest External Patch | Manual RV - External Back Patch | RV Patch - Left Chest External Patch | RV Patch - Back External Patch | LV Patch - Manual RV | LV Patch - RV Patch
Cost | 3 | 6 | 6 | 7 | 7 | 7 | 7 | 6 | 6
Due Date | 1 | 6 | 8 | 8 | 8 | 10 | 10 | 8 | 10
Added Surgical Time for Extra Procedure | 6 | 6 | 5 | 9 | 9 | 8 | 8 | 4 | 3
Ease of Electrode Placement | 8 | 8 | 8 | 9 | 9 | 9 | 9 | 6 | 6
Surgeon Preference | 7 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10
Minimize Added Risk to Patient | 10 | 7 | 6 | 9 | 9 | 8 | 8 | 5 | 4
Reduce Fibrillation Duration | 5 | 9 | 10 | 9 | 9 | 10 | 10 | 9 | 10
Patient Aesthetics | 4 | 10 | 10 | 8 | 8 | 8 | 8 | 10 | 10
Minimized Delivered Energy for Defibrillation | 10 | 10 | 10 | 3 | 3 | 3 | 3 | 10 | 10
Maximizes Myocardial Mass between electrodes | 5 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10
Minimize Post-Op Patient Discomfort | 1 | 10 | 10 | 6 | 6 | 6 | 6 | 10 | 10
TOTALS |  | 509 | 500 | 478 | 478 | 469 | 469 | 463 | 454

## Slide 17: RV Manual Electrode + Left Thoracic Cavity Electrode
RV Manual Electrode
  Right ventricle easily accessible through main incision
  Electrode manually positioned by surgeon using insulated handle
Left Thoracic Cavity Electrode
  Internal chest wall visible through main incision
  Insulated, inflatable shell sutured onto internal chest wall
  Electrode attached to opposite end moved into contact with left ventricle upon inflation of shell

## Slide 18: Details of Design
Capacitor
Inductor
H-Bridge
QUICK-COMBO Therapy Cable
Our Accessory
Patient

## Slide 19: Design Schedule
November 1 – Present design idea to Dr. Damiano
November 3 – Complete specific details of chosen design
November 10 – Finalize analysis needed for design
November 17 – Determine manufacturers of device materials
November 20 – Perform cost analysis for device
November 30 – Complete rough draft of final report
December 7 – Final report due

## Slide 20: Team Responsibilities
Stephen Browning
  Research of patents and current technology
  Development of circuit models
  Development of mathematical models for defibrillation
Kathleen Kafka
  Research of defibrillation techniques
  DesignSafe
  Update group website
  Weekly reports
Kelly Moor
  Research of surgical techniques
  Computer models of electrical systems
  CAD drawings

## Slide 21: Questions?